CN212037509U Epilepsy detection watch, epilepsy detection ring and epilepsy detection device

Data originating from sources other than the EPO may not be accurate, complete, or up to date.

The wording below is an initial machine translation of the original publication. To generate a version using the latest translation technology, go to the original language text and use Patent Translate.

Traducción documento técnico https://chatgpt.com/share/69d056d1-2870-83e9-abac-59124de4a46d

<span class="search">Epilepsy</span> <span class="search">detection</span> watch, <span class="search">epilepsy</span> <span class="search">detection</span> ring and <span class="search">epilepsy</span> <span class="search">detection</span> device

Technical field

This application relates to the technical field of electronic equipment, and in particular to an epilepsy detection watch, an epilepsy detection finger ring and an epilepsy detection device.

Background technique

Epilepsy is the second largest neurological disease in the nervous system after headache. The incidence of epilepsy in our country is 7 per 10,000. There are about 9 million patients with epilepsy. Its pathogenesis is usually due to abnormalities between neurons in the brain nervous system. Cerebral dysfunction diseases caused by electric discharges often manifest as motor twitches, sensory or consciousness disorders. At present, epileptic seizures are detected by means of human motion posture or heart rate change characteristics. For example, there is a motion detection watch on the market, which can realize a comprehensive judgment on the posture of the human body, and guess whether the human body suddenly falls and twitches. However, the above method is prone to false alarms.

Utility model content

The purpose of the embodiments of the present application is to provide an epilepsy detection watch, an epilepsy detection finger ring, and an epilepsy detection device to improve the problem of "currently detecting epileptic seizures by means of human motion posture or heart rate change characteristics, which is prone to false alarms".

This application is implemented like this:

In a first aspect, an embodiment of the present application provides an epilepsy detection watch, which includes: a watch body; a watch strap; connected to the watch body; a first skin inductance sensor arranged on the inner side of the watch strap, when the user wears the epilepsy When detecting the watch, the first skin inductance sensor is in contact with the user's arm, and the first skin inductance sensor is used to detect the human skin electrical signal; the controller; is arranged inside the watch body, and the controller and The first skin inductance sensor is electrically connected; the controller is configured to determine whether the user has epilepsy symptoms based on the human skin electrical signal; and a first communication module; arranged inside the body, the first communication The module is electrically connected to the controller; the first communication module is used to receive the first trigger signal sent by the controller, and send an alarm signal to the remote terminal.

The embodiment of the application provides an epilepsy detection watch provided with a first skin inductance sensor. The first skin inductance sensor can be used to detect human skin signals, where the human skin signals include the tension of human muscles and bioelectric changes. . Therefore, the detection of the electrical skin signal of the human body through the first skin inductance sensor can effectively detect whether the user has an epileptic seizure. Compared with the prior art, the detection rate of the epileptic seizure is improved and misjudgment is reduced.

In combination with the technical solutions provided in the first aspect, in some possible implementation manners, the epilepsy detection watch further includes: a skin inductance chip; the skin inductance chip is arranged inside the watch body; the skin inductance chip is connected to The controller and the first skin inductance sensor are electrically connected; the skin inductance chip is used to receive the human skin signal sent by the first skin inductance sensor, and process the human skin signal, and The processed human skin electrical signal is sent to the controller.

In the embodiment of the present application, the skin inductance chip can collect and process the human skin signal detected by the first skin inductance sensor, and then send the processed human skin signal to the controller, which reduces the controller's processing Pressure improves the stability of operation.

In combination with the technical solution provided by the first aspect, in some possible implementations, the epilepsy detection watch further includes: a heart rate detection module; the heart rate detection module is arranged on the watch body; the heart rate detection module is connected to the watch body; The controller is electrically connected; the heart rate detection module is used to detect the heart rate signal of the human body and send the heart rate signal to the controller.

In the embodiment of the present application, the first skin inductance sensor and the heart rate detection module are simultaneously provided on an epilepsy detection watch, so that the controller can be based on the human skin electrical signal detected by the first skin inductance sensor and the human heart rate signal detected by the heart rate detection module To jointly assist in the detection of the user's epileptic seizure characterization, which further improves the seizure detection rate and reduces misjudgments.

In combination with the technical solutions provided in the first aspect, in some possible implementations, the heart rate detection module includes a light emitting diode and a photosensitive sensor; the light emitting diode and the photosensitive sensor are arranged at the bottom of the watch body; When the user wears the epilepsy detection watch, the light-emitting diode and the photosensitive sensor are in contact with the user's arm; the photosensitive sensor is electrically connected with the light-emitting diode and the controller, respectively; the photosensitive sensor is used for To detect the light signal of the human skin under the illumination of the light emitting diode; and send the light signal to the controller.

In the embodiment of the present application, the photosensitive sensor is used to detect the light signal of the human skin under the illumination of the light-emitting diode, and the controller can obtain the human heart rate change based on the light signal. The heart rate change obtained by this method is accurate and the cost is low.

In combination with the technical solutions provided in the first aspect, in some possible implementations, the heart rate detection module further includes a heart rate detection chip; the heart rate detection chip is arranged inside the watch body; the heart rate detection chip is connected to the The photosensitive sensor and the controller are electrically connected; the heart rate detection chip is used to obtain the heart rate signal based on the light signal sent by the photosensitive sensor; and send the heart rate signal to the controller.

In the embodiment of the present application, the optical signal detected by the heart rate detection module can be processed by the heart rate detection chip, and then the heart rate signal obtained after processing is sent to the controller, which reduces the processing pressure of the controller and improves the stability of operation Sex.

In combination with the technical solutions provided in the first aspect, in some possible implementations, the epilepsy detection watch further includes: a motion detection module, the motion detection module is arranged on the watch body; the motion detection module and the The controller is electrically connected; the motion detection module is used to detect the motion signal of the human body and send the motion signal to the controller.

In the embodiment of the present application, the first skin inductance sensor and the motion detection module are simultaneously provided on an epilepsy detection watch, so that the controller can be based on the human skin electrical signal detected by the first skin inductance sensor and the human motion detected by the motion detection module The signal is used to jointly assist in the detection of the user's epileptic seizure characterization, which further improves the seizure detection rate and reduces misjudgments.

In combination with the technical solution provided by the first aspect, in some possible implementation manners, the epilepsy detection watch further includes: a vibration motor; the vibration motor is arranged inside the watch body, and the vibration motor and the controller Electrically connected; the vibration motor is used to drive the watch body to vibrate according to the second trigger signal sent by the controller.

In the embodiment of the present application, by setting a vibrating motor on the epilepsy detection watch, it can prompt the user that there is a risk of epilepsy at the first time.

In combination with the technical solution provided by the first aspect, in some possible implementation manners, the epilepsy detection watch further includes: a display screen; the display screen is arranged on the top of the watch body; the display screen and the control The device is electrically connected; the display screen is used to display human skin inductance information.

In the embodiment of the present application, the epilepsy detection watch is provided with a display screen, which can display related information such as the human skin electrical signal in real time, so that the user can grasp his own situation in time.

In the second aspect, an embodiment of the present application provides an epilepsy detection finger ring, including a ring body, a second skin inductance sensor, and a second communication module; the second communication module is disposed inside the ring body; the second skin inductance The sensor is arranged inside the ring body; when the user wears the epilepsy detection ring, the second skin inductance sensor is in contact with the user's finger; the second communication module is electrically connected to the second skin inductance sensor The second communication module is used to send the human skin electrical signal detected by the second skin inductance sensor to a remote terminal.

In a third aspect, an embodiment of the present application provides an epilepsy detection device, including the epilepsy detection watch provided in the above first aspect and the epilepsy detection ring provided in the above second aspect; the second communication module and the second communication module A communication module is connected in communication; the second communication module is used to send the human skin electrical signal sent by the second skin inductance sensor to the controller through the first communication module.

Since epilepsy may have half-body seizures, for example, the user's left half of the body may have twitches. Therefore, in an embodiment of the present application, an epilepsy detection device is provided, including an epilepsy detection watch and an epilepsy detection ring. When using the device, the user can wear the epilepsy detection watch and the epilepsy detection ring on different hands, thereby realizing bilateral detection of the human body, and improving the detection rate of epileptic seizures compared with the prior art.

Description of the drawings

In order to explain the technical solutions of the embodiments of the present application more clearly, the following will briefly introduce the drawings that need to be used in the embodiments of the present application. It should be understood that the following drawings only show some embodiments of the present application. Therefore, It should not be regarded as a limitation of the scope. For those of ordinary skill in the art, other related drawings can be obtained from these drawings without creative work.

FIG. 1 is a schematic structural diagram of an epilepsy detection watch from a first perspective according to an embodiment of the application.

FIG. 2 is a block diagram of modules of an epilepsy detection watch provided by an embodiment of the application.

FIG. 3 is a schematic structural diagram of an epilepsy detection watch from a second perspective according to an embodiment of the application.

FIG. 4 is a schematic structural diagram of an epilepsy detection watch from a third perspective according to an embodiment of the application.

FIG. 5 is a schematic structural diagram of an epilepsy detection ring provided by an embodiment of the application.

Icon: 100- epilepsy detection watch; 10- body; 11- controller; 12- first skin inductance sensor; 121- skin inductance chip; 13- first communication module; 14-PCB motherboard; 15- heart rate detection module; 151-light emitting diode; 152-photosensitive sensor; 153-heart rate detection chip; 16-motion detection module; 17-vibration motor; 18-display screen; 20-strap; 200-epilepsy detection ring; 30-ring body; 31- The second skin inductance sensor; 32-the second communication module.

Detailed ways

The technical solutions in the embodiments of the present application will be clearly and completely described below in conjunction with the drawings in the embodiments of the present application.

At present, epileptic seizures are detected by means of human motion posture or heart rate change characteristics. For example, there is a motion detection watch on the market, which can realize a comprehensive judgment on the posture of the human body, and guess whether the human body suddenly falls and twitches. However, the above method is prone to false alarms.

In view of the above-mentioned problems, the inventor of the present application, after research and exploration, proposes the following embodiments to solve the above-mentioned problems.

Please refer to FIG. 1 to FIG. 2. An embodiment of the present application provides an epilepsy detection watch 100. It includes: a watch body 10, a watch band 20, a controller 11, a first skin inductance sensor 12, and a first communication module 13.

The controller 11 is electrically connected to the first skin inductance sensor 12 and the first communication module 13 respectively.

The watch body 10 is connected to the watch band 20, and the shape of the watch body 10 may be square or circular, which is not limited in this application. The first skin inductance sensor 12 is arranged on the inner side of the watch band 20. The inner side of the watchband 20 is the side close to the user's arm. When the user wears the epilepsy detection watch 100, the user's arm is in contact with the inner side of the watchband 20. By arranging the first skin inductance sensor 12 on the inner side of the watchband 20, when the user wears the epilepsy detection watch 100, the first skin inductance sensor 12 is in contact with the user's arm, which can facilitate the first skin inductance sensor 12 to touch the human skin. Detection of electrical signals. The electrical skin signal of the human body includes the tension of human muscles and the changes of bioelectricity. The first skin inductance sensor 12 sends the human skin electrical signal to the controller 11 after detecting the human skin electrical signal.

Wherein, the controller 11 is arranged inside the meter body 10, and a PCB (Printed Circuit Board) main board 14 can be arranged inside the meter body, and circuit modules such as the controller 11 can be integrated on the PCB main board 14. The controller 11 is used to determine whether the user has epilepsy symptoms based on the human skin electrical signal, and when it is determined that the user has epilepsy symptoms, send a first trigger signal to the first communication module, so that the first communication module sends an alarm signal to a remote terminal .

The aforementioned controller 11 may include multiple processing units, and the controller 11 is an integrated circuit chip with signal processing capabilities. For example, the controller 11 may be an industrial computer or a single-chip computer. The controller 11 may also be a general-purpose processor, for example, it may be a digital signal processor (DSP), an application specific integrated circuit (ASIC), a discrete gate or transistor logic device, or a discrete hardware component. Or execute the methods, steps, and logical block diagrams disclosed in the embodiments of the present application. In addition, the general-purpose processor may be a microprocessor or any conventional processor or the like. In the embodiment of the present application, the model of the controller 11 is DA14683.

The first communication module 13 is also arranged inside the watch body 10. The first communication module 13 is configured to receive the first trigger signal sent by the controller 11 and send an alarm signal to the remote terminal. Among them, the first communication module 13 may be a 3G, 4G module, a Bluetooth module, or a ZigBee module. The remote terminal can be a user's mobile phone, tablet computer or other equipment. For example, if the user A wears the epilepsy detection watch 100 described above, the mobile phone of the spouse of the user A or the cell phone of the child of the user A may be the mobile phone of the spouse of the user A or the cell phone of the child of the user A that is communicatively connected with the epilepsy detection watch 100.

In summary, the embodiment of the present application provides an epilepsy detection watch 100 provided with a first skin inductance sensor 12. The first skin inductance sensor 12 can be used to detect human skin signals, where the human skin signals include human body Muscle tension and bioelectric changes. Therefore, detecting the electrical skin signal of the human body through the first skin inductance sensor 12 can effectively detect whether the user has an epileptic seizure. Compared with the prior art, the seizure detection rate is improved and misjudgment is reduced.

Referring to FIG. 3, optionally, the epilepsy detection watch 100 described above further includes a skin inductance chip 121. The skin inductance chip 121 is arranged inside the body.

The skin inductance chip 121 is electrically connected to the controller 11 and the first skin inductance sensor 12 respectively. The skin inductance chip 121 is used to receive the human skin electrical signal sent by the first skin inductance sensor 12, process the human skin electrical signal, and send the processed human skin electrical signal to the controller 11.

Optionally, in the embodiment of the present application, the model of the skin inductor chip 121 is MAX32600.

In the embodiment of the present application, the skin inductance chip 121 can collect and process the human skin signal detected by the first skin inductance sensor, and then send the processed human skin signal to the controller 11, which reduces The processing pressure of 11 improves the stability of operation.

However, it is found in the research that there is still contingency in the single detection method, that is, the single detection method still has the possibility of misjudgment. Therefore, please refer to FIG. 4. Optionally, the epilepsy detection watch 100 further includes a heart rate detection module 15. The heart rate detection module 15 is arranged on the watch body 10. The heart rate detection module 15 is electrically connected to the controller 11. The heart rate detection module 15 is used to detect the heart rate signal of the human body and send the heart rate signal to the controller 11.

In the embodiment of the present application, the first skin inductance sensor 12 and the heart rate detection module 15 are simultaneously provided on an epilepsy detection watch 100, so that the controller 11 can detect the human skin electrical signal and the heart rate detected by the first skin inductance sensor 12 The human heart rate signal detected by the module 15 is used to jointly assist in the detection of the user's epileptic seizure characterization, which further improves the seizure detection rate and reduces misjudgments.

Specifically, in the embodiment of the present application, the heart rate detection module 15 includes a light emitting diode 151 and a photosensitive sensor 152. The light emitting diode 151 and the photosensitive sensor 152 are arranged at the bottom of the watch body 10, and the bottom of the watch body 10 is also close to the end of the user's arm. When the user wears the epilepsy detection watch 100, the light emitting diode 151 and the photosensitive sensor 152 are in contact with the user's arm.

Among them, the photosensitive sensor 152 is electrically connected to the light emitting diode 151 and the controller 11 respectively. The photosensitive sensor 152 is used to detect the light signal of the human skin under the illumination of the light emitting diode 151; and send the light signal to the controller 11. The controller 11 can obtain the human heart rate change based on the light signal. The heart rate change obtained by this method is accurate and the cost is low.

Optionally, the heart rate detection module 15 further includes a heart rate detection chip 153 (as shown in FIG. 1). The heart rate detection chip 153 is arranged inside the watch body 10. The heart rate detection chip 153 is electrically connected to the photosensitive sensor 152 and the controller 11 respectively. The heart rate detection chip 153 is configured to obtain the heart rate signal based on the light signal sent by the photosensitive sensor 152; and send the heart rate signal to the controller 11.

In the embodiment of the present application, the model of the heart rate detection chip 153 is YK1303P.

In other embodiments, the aforementioned heart rate detection module 15 may also be a heart rate sensor, such as LST1303. In this regard, this application is not limited.

In the embodiment of the present application, the light signal detected by the heart rate detection module 15 can be processed by the heart rate detection chip 153, and then the heart rate signal obtained after processing is sent to the controller 11, which reduces the processing pressure of the controller 11 and improves Improve the stability of operation.

Please continue to refer to FIGS. 1 and 2, as another way to further improve the detection rate of epileptic seizures, the epilepsy detection watch 100 further includes a motion detection module 16. The motion detection module 16 is provided on the watch body 10. The motion detection module 16 is electrically connected to the controller 11. The motion detection module 16 is used to detect the motion signal of the human body and send the motion signal to the controller 11.

Optionally, in the embodiment of the present application, the motion detection module 16 is a six-axis motion sensor, and its specific model may be LSM6DS3. The sensor integrates a three-axis accelerometer and a three-axis gyroscope to evaluate the physiological information of the human body movement, and can monitor the physiological characteristic signals of the seizure movement during the seizure. Of course, in other implementations, the motion detection module 16 may also only include a three-axis accelerometer or a three-axis gyroscope, and the model of the six-axis motion sensor is not limited.

In the embodiment of the present application, the first skin inductance sensor 12 and the motion detection module 16 are simultaneously provided on an epilepsy detection watch, so that the controller 11 can be based on the human skin electrical signal detected by the first skin inductance sensor 12 and the motion detection module 16 The detected human motion signals are used to jointly assist in the detection of the user's epileptic seizure characterization, which further improves the seizure detection rate and reduces misjudgments.

It is understandable that the epilepsy detection watch 100 provided in the embodiment of the present application may also be provided with the first skin inductance sensor 12, the heart rate detection module 15 and the motion detection module 16 at the same time. Through the multi-modal information of heart rate signal, human body motion signal, and human skin electrical signal, the accuracy of epileptic seizure prediction is further improved.

Optionally, the epilepsy detection watch 100 further includes a vibration motor 17. The vibration motor 17 is arranged inside the watch body 10, and the vibration motor 17 is electrically connected to the controller 11. The vibration motor 17 is used to drive the watch body 10 to vibrate according to the second trigger signal sent by the controller 11. In the embodiment of the present application, by setting the vibration motor 17 on the epilepsy detection watch 100, the user can be alerted to the danger of epilepsy at the first time.

Optionally, the epilepsy detection watch 100 further includes a display screen 18. The display screen 18 is arranged on the top of the watch body 10. The display screen 18 is electrically connected to the controller 11, and the display screen 18 is used to display human skin inductance information. Of course, if the epilepsy detection watch 100 is also provided with the heart rate detection module 15 and the motion detection module 16, the display screen 18 is also used to display the heart rate signal and the human motion signal. By providing a display screen on the epilepsy detection watch 100, relevant information such as human skin electrical signals can be displayed in real time, so that the user can grasp his own situation in time.

Wherein, the display screen 18 may be an LED (light emitting diode) touch screen. The touch screen can be convenient for users to operate by hand.

Of course, the epilepsy detection watch 100 provided in the embodiment of the present application further includes a power module. The power supply module is used to supply power to the electrical structures in the epilepsy detection watch 100. For example, the power module is electrically connected to the first skin inductance sensor 12, the controller 11, and the first communication module 13 respectively. For another example, the power module is also electrically connected to the skin inductor chip 121. For another example, the power module is also separately powered by the heart rate detection module 15, the motion detection module 16, the vibration motor 17, and the display screen 18.

In order to improve the stability of the power supply, a power supply regulator capacitor can also be connected to the power supply module side. The power supply module may also be provided with a power management chip, through which different devices can be provided with different working voltages.

The aforementioned power module may be a lithium battery. Optionally, a charging interface is further provided on one side of the watch body 10, and the charging interface is electrically connected to the lithium battery. Correspondingly, the charging port may not be provided on one side of the watch body 10, and when the epilepsy detection watch 100 is out of power, only the battery needs to be replaced.

Of course, the power module may also include a wireless induction charging coil. Through the wireless induction charging coil, the epilepsy detection watch 100 can have a wireless charging function.

Hereinafter, the use of the epilepsy detection watch 100 described above will be specifically described. When using the epilepsy detection watch 100, the user can wear the epilepsy detection watch 100 on any hand. The first skin inductance sensor 12, the heart rate detection module 15, and the motion detection module 16 on the epilepsy detection watch 100 can be detected at regular intervals, for example, every 15 seconds, or of course can be detected in real time. . When the epilepsy detection watch 100 detects an abnormality in the user, for example, after the detection signals of the first skin inductance sensor 12, the heart rate detection module 15, and the motion detection module 16 exceed the warning value, first, the controller 11 in the epilepsy detection watch 100 will The vibration motor 17 sends a second trigger signal to drive the watch body 10 to vibrate, thereby prompting the user to be at risk of seizures. Then, the controller 11 sends a first trigger signal to the first communication module 13 so that the first communication module 13 sends an alarm signal to the remote terminal.

Based on the foregoing use process, the embodiment of the present application also provides a confirmation mechanism. That is, after the vibrating motor 17 drives the watch body 10 to vibrate, the user can confirm whether there is an epilepsy symptom. If the epilepsy detection watch 100 makes a misjudgment, the user can release the alarm by short pressing the switch button 3 times. Through this epilepsy detection confirmation mechanism, the epilepsy detection watch 100 can realize the self-learning update of epileptic seizures and pre-prediction algorithms, and update the updated learning model to the memory of the DA14683 chip to improve the epilepsy detection watch 100 The accuracy of detection also realizes the personalized customization function for patients.

Referring to FIG. 5, based on the same concept, an embodiment of the present application provides an epilepsy detection finger ring 200, which includes a ring body 30, a second skin inductance sensor 31 and a second communication module 32. The second communication module 32 is arranged inside the ring body. The second skin inductance sensor 31 is arranged inside the ring body. When the user wears the epilepsy detection finger ring 200, the second skin inductance sensor 31 contacts the user's finger.

The second communication module 32 is electrically connected to the second skin inductance sensor 31. The second communication module 32 is configured to send the human skin electrical signal detected by the second skin inductance sensor 31 to the remote terminal.

It should be noted that the epilepsy detection finger ring 200 provided in this embodiment is not provided with a controller. Therefore, the detection of epileptic seizures requires the help of remote terminals, such as the user's mobile phone, tablet computer, etc. The epilepsy detection ring 200 is used for the user's mobile phone of the detected human skin electrical signal, and the user's mobile phone can be used to determine whether the user has epilepsy symptoms based on the human skin electrical signal.

Based on the same concept, an embodiment of the present application provides an epilepsy detection device. It includes the epilepsy detection watch 100 and the epilepsy detection ring 200 provided in the foregoing embodiment.

The second communication module 32 is in communication connection with the first communication module 13. The second communication module 32 is configured to send the human skin electrical signal sent by the second skin inductance sensor 31 to the controller 11 through the first communication module 13.

Since epilepsy may have a half-body seizure, for example, a user's left body may have twitches, in an embodiment of the present application, an epilepsy detection device is provided, which includes an epilepsy detection watch 100 and an epilepsy detection ring 200. When using the device, the user can wear the epilepsy detection watch 100 and the epilepsy detection ring 200 on different hands, thereby realizing bilateral detection of the human body, and improving the detection rate of epileptic seizures compared with the prior art.

In the description of this application, it should be noted that the orientation or positional relationship indicated by the terms "inner" and "outer" is based on the orientation or positional relationship shown in the drawings, or is usually placed when the application product is used. The orientation or positional relationship is only for the convenience of describing the application and simplifying the description, rather than indicating or implying that the device or element referred to must have a specific orientation, be constructed and operated in a specific orientation, and therefore cannot be understood as a limitation of the application. In addition, the terms "first", "second", etc. are only used for distinguishing description, and cannot be understood as indicating or implying relative importance.

It should also be noted that, unless otherwise clearly specified and limited, the terms "setting" and "connection" should be understood in a broad sense. For example, it can be a fixed connection, a detachable connection, or an integral connection; it can be a direct connection. Connected can also be indirectly connected through an intermediate medium, which can be the internal communication between two components. For those of ordinary skill in the art, the specific meaning of the above-mentioned terms in this application can be understood under specific circumstances.

The foregoing descriptions are only examples of the present application and are not used to limit the protection scope of the present application. For those skilled in the art, the present application may have various modifications and changes. Any modification, equivalent replacement, improvement, etc., made within the spirit and principle of this application shall be included in the protection scope of this application.